# Requirement for monitoring the C-TFN with TimeMap

Note this version of the document relates to the Pathfinder implementation of TimeMap.

GEANT is building a C-TFN network to support the European metrology community time/frequency distribution needs.  For details see the report here: Core-TFN-Incubator-Final-Report

This document discusses the high-level requirements for C-TFN time/frequency monitoring using TimeMap.  In particular, it lists the equipment to be monitored, existing management software, what is missing from these existing tools, and how we can build a single user-interface to meet the metrology community needs.

In this study we consider TimeMap to be a monitoring tool and it should not be used for equipment configuration.

## Users

TimeMap should provide a user interface which allows the following users to monitor the C-TFN network:

- GEANT Operations Centre

- All participating National Metrology Institutes (NMIs)

- All participating National Research and Education Networks (NRENs)

- Any end-users of time or frequency services, for example ESA.

## User actions

Each of the users will need to carry out specific monitoring actions on the C-TFN, these are summarized here.

GEANT Operations Centre (OC) actions:

The following actions are performed by the GEANT Operations Centre for the purpose of first-line maintenance of the C-TFN network.

In this table EM refers to the local network Element Manager (EM) or local CLI interface for equipment configuration. This EM is normally provided by the equipment vendor.

(this section to be reviewed and updated by GEANT, Guy Roberts to lead this work)

| Action | Vendor EM/CLI | TimeMap |
| --- | --- | --- |
| Aggregated view of up/down status of all amplifiers in the C-TFN. This will allow the OC to rapidly become aware of hardware faults of bi-directional EDFA amplifiers | No – On a per-node basis only | Yes (use ILA status traps) |
| Aggregated view of up/down status of the state of frequency services | no | Yes |
| Aggregated view of up/down status of the state of time services – both ELSTAB and white rabbit | no | Yes |
| Monitor the amplifiers for indications of faults. Carry out an investigation of root-cause. | yes | No |
| Aggregated view fibre links for fibre cuts and other forms of fibre degradation | no | Yes (use LoS traps) |
| Monitor and adjust optical amplifier gain if needed when adding new services | yes | no |
| Monitor optical power levels to support NMIs during service turn-up | yes | no |
| Monitor EDFA amplifiers configuration to check if excessive double Rayleigh scattering can occur (this can affect transmission quality) Is this too complex for the GEANT OC, for NMI specialists to use this data? | no (how to predict this, using calculations?) | no |
| Support the NMIs in turning up new time services. |  |  |
| Support the NMIs in turning up new frequency services |  |  |

NMI frequency services actions:

The following actions are performed by the NMI for the purpose of maintenance of their local flywheel.

(this section to be reviewed and updated by NMIs, Domenico V. to lead this work via SIG-TFN)

| Action | Vendor EM | TimeMap |
| --- | --- | --- |
| Turn up new frequency services using the RLS equipment | Yes | No |
| Monitor RLS performance and adjustment of any parameters to optimize performance. | Yes | No |
| Monitor EDFA performance and adjustment of any parameters to optimize performance. Note: the NMIs will not have access to the CLI of the EDFAs, hence the need for these parameters in TimeMap. Tuning of the gain will need to be done by the GEANT OC (or NREN operator of the fibre link.) | No | Yes |
| Monitor EDFA amplifiers configuration to check if excessive double Rayleigh scattering can occur (this can affect transmission quality) Is this too complex for the GEANT OC, for NMI specialists to use this data? | no (how to predict this, using calculations?) | no |
| Monitor the status of the flywheel (free running, steered). How is the local oscillator that is being transmitted stabilized? - Locked to the local maser cavity - Steered by the local time reference (Optical reference) Need further information from NMIs about how this information is shared | ? | ? |
| Detect frequency cycle slips in the RLS due to faults such as reflections on the fibre or excess phase shifts due to fibre movement. Mark time/frequency data as possible cycle slips Note: the information of cycle slips will depend on the configuration of the link, if MLS with up/down link is implemented then cycle slips can be detected. Otherwise this is a probable event information only. | ? | ? |
| Monitor RLS/MLS equipment for hardware faults, replacing faulty hardware. | yes | ? |
| Early warning of RLS hardware or other major faults | no | yes |

NMI time services actions:

The following actions are performed by the NMI for the purpose of maintenance of their White Rabbit switches or ELSTAB equipment.

(this section to be reviewed and updated by NMIs, Domenico V. to lead this work via SIG-TFN)

| Action | Vendor EM | TimeMap |
| --- | --- | --- |
| Turn up new WR or ELSTAB services | Yes | No |
| Calibrate WR or ELSTAB equipment | Yes | No |
| Monitor WR or ELSTAB equipment for hardware faults, replacing faulty hardware | Yes | No |
| Early warning of WR or ELSTAB fault for rapid response. | No | Yes |
| Perform software updates of WR boxes | Yes | No |

End user for frequency actions:

The following actions are performed by the end-users (such as CERN) for the purpose checking the offset of their frequency service in relation to a national frequency reference.

(this section to be reviewed and updated by end-users, Domenico V. to lead this work via SIG-TFN)

| Action | Vendor EM | TimeMap |
| --- | --- | --- |
| Login to TimeMap (or NMI database) and collect comparator frequency offset measurement data. | no | yes |

End user for time actions:

The following actions are performed by the end-users (such as ESA) for the purpose checking the offset of their WR (or ELSTAB) time feed in relation to a national time reference.

(this section to be reviewed and updated by end-users, Domenico V. to lead this work via SIG-TFN)

| Action | Vendor EM | TimeMap |
| --- | --- | --- |
| Login to TimeMap (or NMI database) and collect comparator time offset measurement data. |  | yes |
| Check up/down status of WR (or ELSTAB) service |  | yes |

## T/F equipment and existing management tools

The following tables summarize the hardware that makes up the C-TFN, and for each type of equipment the software and its monitorable parameters are listed.

EDFA Bidirectional amplifier

The bidirectional in-line amplifier (ILA) increases the optical power of the light transiting through the amplifier.  The operation is bi-directional, i.e. light is amplified in both directions.

- TimeMap to report hardware up/down status.

- No data needed, TimeMap to provide URL to vendor’s local tool for everything apart from up/down status.

(this section to be reviewed and updated by PSNC)

| Equipment | Monitored Parameters | Monitor in TimeMap? |
| --- | --- | --- |

| Bi-dir EDFA PIKTime PSNC have developed a monitoring system for PIKTIME amplifiers. This allows for full OAMP | SNMP Traps: laser bias out of range, LoS (to be confirmed), any H/W fault alarm, LoS of management channel. | Future TimeMap (further work required to understand the user requirements for the traps to be reported) |
| --- | --- | --- |
|  | Gain - we can monitor the set gain level on each amplifier (in our case we have stabilized gain). In general, you could monitor the pump current, however, this information is only useful to experts. This data can be used for link optimization. | No |
|  | Input and output power levels (4 parameters) can be monitored. Unfortunately, it is impossible to distinguish between OC signal and ASE noise in amplifiers. So it is necessary to read them with understanding | no |
|  | We also have the ability to monitor separately the level of the modulated signal and several me important parameters | no |
| Other vendor? | None for pathfinder | no |

Regenerator Laser Station

The regenerator laser station is used to remove phase noise from the frequency signal after it has passed over the optical fibre.  The pathfinder will use the RLS equipment from PIC/PSNC.  See the description of an RLS from Exail: https://www.muquans.com/product/regeneration-station/

- TimeMap to report RLS hardware up/down status.

- No monitoring data needed needs to be reported in TimeMap, a URL to vendor’s local tool is sufficient for this.

(this section to be reviewed and updated by PSNC)

| Equipment | Monitored Parameters | Monitor in TimeMap? |
| --- | --- | --- |
| RLS PSNC/ PIKTIME PIKTIME management system. management system. Supports full OAMP | States of individual inputs: locked/unlocked (both uplink and downlink) This parameter provides information on whether a given link has correctly locked. The unlocked state indicates that the link “does not work” and the other values are unreliable. | Yes available to TimeMap now Link state: locked/unlocked Link state for both uplink and downlink. |
|  | Beat power, i.e. the received optical power on individual inputs (both uplink and downlink) | Yes available to TimeMap now Beat power measurement live in dBm |
|  | On the uplink port we have two noise parameters Low and High which tell us the amount of noise of a given type. These parameters are helpful in optimizing the gain of individual optical amplifiers. | Yes available to TimeMap now High frequency noise Low frequency noise Collected every 10 seconds |
|  | On the DownLink ports we currently have monitoring of the jitter parameter. Currently, with Kuba we are working on additional parameters which, together with jitter, would inform us about the possibility of cycleslips on particular transmission directions Local->Remote or Remote->Local | Yes available to TimeMap now Jitter Max value during a 10 second window. |
|  | other parameters are probably less crucial for monitoring the correctness of link operation | No |
|  | The PIKTIME devices optimize polarization continuously and therefore there are no interruptions in OC transfer related to polarization optimization are expected in normal operation. In the case of excess polarization drift the beat power will drop and PIKTIME equipment will automatically re-align the polarization. This re-alignment process can be signal affecting. | No Need to discuss with the NMIs if there is a need to report the polarization information at the receiver. |
| Other | No other RLS vendors for the pathfinder It is believed that Exail equipment makes this polarization optimization however it involves signal unavailability? If so, the monitoring system should also consider such a situation and mark a certain range of measurements as unreliable/false? |  |

White Rabbit Switch

A White Rabbit Switch is an Ethernet switch with Precision Time Protocol (PTP) enabled and some additional calibration capabilities that increase the accuracy of time stamped packets.

- WR switch is not used in the pathfinder

- TimeMap to report hardware up/down status (future feature).

- No monitoring needed, provide URL to cli if needed for maintenance.

(this section to be reviewed and updated by TBC)

| Equipment | Monitoring/ configuration Parameters | Monitor in TimeMap? |
| --- | --- | --- |
| WR switch Vendor? | cli for monitoring and configuration | No |
|  | Ethernet swich configuration | No |
|  | Time calibration parameters | No |
|  | SNMP trap for switch status? | Yes (future) |

ELSTAB

ELSTAB is an alternative to White Rabbit and has been developed by PSNC/PIKTIME.  It is used to distribute time based on a modulated optical wavelength.

- TimeMap to report ELSTAB hardware up/down status.

- No data needed, provide URL to vendor’s local tool.

(this section to be reviewed and updated by PSNC)

| Equipment | Monitoring/ configuration Parameters | Monitor in TimeMap? |
| --- | --- | --- |
| PIKTIME ELSTAB equipment | Up/down status of hardware | yes |
|  | ? |  |
|  | ? |  |

Frequency comb/comparator

This equipment is made up of two parts, the frequency comb and the frequency counter.  The comb is passive and does not have any monitoring information relevant to TimeMap.  The frequency counter measures the difference in frequency between the local reference and the incoming signal from the fibre.  For the pathfinder the following counter is used:  http://www.kplusk-messtechnik.de/products/fxe_19.htm

- No status or monitoring information for comb

- TimeMap to report the hardware up/down status of the comparator.

- The frequency offset data set is required to be reported in TimeMap.

(this section to be reviewed and updated by an NMI, Guy Roberts and Domenico to reach out to NMIs to collect these requirements and understand how each NMI’s script works)

| Equipment | Parameters | Monitor in TimeMap |
| --- | --- | --- |
| Frequency comparator Vendor K+K http://www.kplusk-messtechnik.de/products/fxe_19.htm Local Windows PC collects the text file. NMI home-developed script is currently used for local data collection. Datasets are archived locally on a server. Is there an opportunity to have each NMI share the same scripts? Or can they give GEANT the script and we can collect raw data and run the script ourselves. | Hardware up/down status? | Future TimeMap |
|  | Frequency offset (difference) data | Future TimeMap |
|  | Phase offset (difference) data | Future TimeMap |

Time interval counter

The time interval counter is used to compare the time difference between two time references.

- TimeMap to report hardware up/down status of the comparator.

- The time offset data set is required to be reported in TimeMap.

(this section to be reviewed and updated by an NMI)

| Equipment | Parameters | Monitor in TimeMap |
| --- | --- | --- |
| time interval counter NMI home-developed script is currently used for local data collection. Datasets are archived locally on a server. | Hardware up/down status? | yes |
|  | Time offset dataset | yes |